Customer Product Quality Requirements Form

Part Number: 1500		    X      New Product       	               Revised Product

Company: OXY Whichita		Contact:  Dale Jones

Address:

Phone:					email:  Dale_Jones@oxy.com

*If not already indicated, please specify if this product needs to be prepared or verified to an

industry standard procedure as ASTM, EPA, APHA or other method:

Note: in the event that the above method becomes revised, it will be the customer’s responsibility to provide updated information / revision requests to Aqua Solutions.

1500		BUFFER SOLUTION  pH 7.00 +- 0.01 @ 25 DEG C. N.I.S.T. TRACEABLE  (YELLOW)

NOTE 1: THIS SOLUTION MUST STIR FOR 1-2 FULL DAYS BEFORE CHECKING THE PH.

USE QCSP20B WHICH USES ACTUAL N.I.S.T. MATERIAL TO CALIBRATE THE METER.

BEFORE THE PRODUCT IS ABLE TO BE RELEASED FOR PACKAGING, IT MUST BE WITHIN SPECIFICATIONS, I.E. +- 0.01 OF  PH FOR 2 CONSECUTIVE TESTS THAT ARE AT LEAST 1 DAY APART. DO NOT RELEASE PRODUCT THAT HAS NOT PASSED TWO CONSECUTIVE TESTS.

NOTE 2: ONCE THE BUFFER HAS BEEN APPROVED FOR PACKAGING, IT MUST BE  PACKAGED  WITHIN 1 DAY AFTER APPROVAL.

NOTE 3: DO NOT ADD COLOR TO THE BUFFER UNTIL THE DAY BEFORE IT WILL BE PACKAGED TO PREVENT STAINING OF THE TANK.

1-LITER FORMULATION

FOR EACH LITER OF FINAL SOLUTION, DISSOLVE 6.8 GRAMS OF POTASSIUM PHOSPHATE MONOBASIC (P5600) AND 2.328 GRAMS OF SODIUM HYDROXIDE 50% W/W  (S2923) (ADJUST BASED ON ACTUAL ASSAY) INTO 900 MLS OF DEIONIZED WATER.(9765). DILUTE TO 1 LITER WITH DEIONIZED WATER (9765) AND STIR FOR AT LEAST 1 DAY BEFORE CHECKING THE  pH. ONCE APPROVED FOR PACKAGING ADD  0.004 GRAMS OF THYMOL BLUE SODIUM SALT (T4954) FOR EACH LITER OF FINAL SOLUTION TO TURN THE BUFFER YELLOW.

REVISION: 2.4  	DATE:  03-26-2019        		 INITIAL:_STN

EXPIRATION: 2 YEARS

FILTERING INSTRUCTIONS:  THIS BATCH IS USUALLY MADE IN A LARGE TANK AND IS FILTERED

THROUGH 75UM INLINE FILTER.  IF THIS BATCH IS MADE IN A SMALLER TANK USE GLASS WOOL TO REMOVE ANY STIR BAR DEBRIS.

SPECIAL PACKAGING: IT MUST BE PACKAGED 1 DAY AFTER APPROVAL

PACKAGING:  CLEAR PLASTIC OR CUBES

HMIS 100B

PERSONAL PROTECTION: GLOVES, GLASSES.

STANDARDIZATION: USE NIST TRACEABILITY PORTION OF QCSP20B TO STANDARDIZE TO 10.00 +- 0.02 @ 20DEG C.

APPEARANCE/ ODOR: CLEAR YELLOW ODORLESS LIQUID

DISPOSAL: GENERAL

REVISION 1.0, 08/10/96: REVISED TO INCLUDE QCSP STANDARDIZATION METHOD AND TO DECREASE      SPECIFICATION WHEN STANDARDIZING TO 7.00 +- 0.01 @25° C

REVISION 2.0 11/09/99 REVISED BATCH SHEET AND BATCH CARD TO EXACT FORMULATION

REVISION: 2.1 10/30/01 REVISED TO PUT THAT THIS SOLUTION MUST STIR FOR 1-2 FULL DAYS AND PASS TWO CONSECUTIVE TEST BEFORE PACKAGING AND THAT IT MUST BE PACKAGED WITHIN 1 DAY AFTER  APPROVAL.AND ADD THE COLOR AFTER YOU DILUTE TO 1 LITER AND RIGHT BEFORE PACKAGING.

REVISION: 2.2, 05/21/03, REVISED THIS PRODUCT TO BE MADE FROM 1501 CLEAR BUFFER WHICH HAS ACTUALLY BEEN  DONE FOR A WHILE IN ORDER TO BE ABLE TO USE 2 DIFFERENT LOT NUMBERS. STN

REVISION 2.3 ,04/27/11 :  ADDED FILTERING INSTRUCTIONS. RK

Revision 2.4, 03-26-2019: Updated description to change specifications to +- 0.01 from +- 0.02. STN

***Please make any revisions, sign and email or fax back to Aqua Solutions @ 281-479-1833

Customer Authorization Signature:  ________________________________________  Date:         /           /

Laboratory Approval: __________________________________________________  Date:       /         /